## Sheet1
DL/R_Type | Summary | Requirement Details
1 | General Requirement
1.1 | Safety | Ventilators must be designed to protect patients from electric shock, mechanical hazards, and excessive outputs of pressure (35 cmH2O) and volume
1.2 | Environmental conditions | The ventilator should function safely within a specified range of environmental conditions, including temperature and humidity
2 | Mode and Operation
2.1 | Operating mode | The ventilator shall allow changing the ventilation mode while ventilation support is inactive or active
2.2 | Start/Stop | The ventilator shall provide switch input to allow start/stop ventilation operation
2.3 | Monitoring | The ventilator shall provide monitoring switch to allow display all essential patient vitals such as tidal volume, pressure, flow, inhale/exhale time
2.4 | Press control | The pressure controller ventilator shall be operate within a range of 5 to 50 cm H2O
2.5 | PEEP adjustment | The ventilation modes that utilizes PEEP adjustment shall range between 5 to 10 cmH2O
2.6 | Breath rate | The ventilation modes that utilizes breath rate shall range between 1 to 60 breath/min
2.7 | Insp time | The Inspirational time adjustment range shall be 0.3s to 2.4s with an I/E ratio not exceeding 1/2 for pressure controlled ventilator
2.8 | Pressure level | The rise time adjustment range shall allow 6 levels. The associated quantitative rise times will adapt to inspiration time and the adjusted inspiratory pressure value
2.9 | Heater wire humidifier | The temperature along the inspiratory limb shall be monitored and controlled at 34degC in the non-invasive mode and 40degC in the invasive mode
2.11 | Humidity requirement | The ventilator shall deliver gas with 50% relative humidity with absolute humidity of 10mg/L at 22degC at nose/mouth
3 | Display
3.1 | Patient monitoring | The ventilator should provide monitoring capabilities for patient parameters such as pressure, volume, and gas concentration
4 | Alarms
4.1 | Operating mode short circuit | Operating mode dial switch logic shall able to detect shorted high or ground faults
4.2 | Pressure fault | Pressure sensor reading shall not fall below 5 or rise above 50 cmH2O while ventilation support is active
4.3 | Occlusion | The ventilator shall initiate a high priority alarm if no significant volume can be delivered to the patient while the target pressure is reached
4.4 | Alarm failure | The ventilator shall initiate a high priority alarm if the primary audio/visual alarm system is not delivering alarm signal